## Slide 1: GP-SANS

## Slide 2: Instrument View From Top in the world of instrument scientists and EPICS
source
sample
Silicon window
detector
SampleToSi
sample_detector_distance (sdd)

## Slide 3: From the world of IS/EPICS to the world of drt-sans
In drt-sans, data from a NeXus file is loaded into a Mantid workspace
In a Mantid workspace, an instrument (object) is created for
  Neutron source position
  Sample position
  Detector pixels position
Instrument geometry is set to a Cartesian coordinate system in drt-sans (Mantid)
  A fixed origin/nominal position
  Sample can be moved
  Detector can be moved
  Source can be moved
  Silicon window CANNOT be moved (i.e., fixed)

## Slide 4: From the world of IS/EPICS to the world of drt-sans (cont)
Meta data in the 3 worlds
Instrument scientists | EPICS | drt-sans
Sample to silicon window distance | SampleToSi | swd
Sample to detector distance | sample_detector_distance | sdd

## Slide 5: drt-sans: Expected Instrument View From Top (+Y)
source
sample
Silicon window
detector
SampleToSi (swd)
sample_detector_distance (sdd)
(0, 0, 0)
+ Z
+ X
Center of silicon window is at nominal position/origin (0, 0, 0)
Sample is at (0, 0, -swd)
Detector is at (0, 0, sdd - swd)
(0, 0, sdd - swd)

## Slide 6: drt-sans: Instrument view from Top in Current Implementation (wrong)
source
sample
Silicon window
detector
sample_detector_distance (sdd)
(0, 0, 0)
+ Z
+ X
Issue: sample always is positioned at (0, 0, 0). Sample to silicon distance is never considered
Is it critical?
Yes if sample to silicon window distance is to be overwritten as new sample sample to detector distance can be calculated wrong
No in other cases as long as sample to detector distance is correct
Cause: Mantid instrument definition file (IDF) does not consider meta data SampleToSi
Solution:
Move the sample toward source by swd to (0, 0, -swd)
Move the detector toward source/sample by swd to (0, 0, sdd - swd)
Such that the sample-detector distance is not changed from the value recorded in EPICS
Then the instrument will be corrected to what is shown in previous slide.
(0, 0, sdd)

## Slide 7: drt-sans: Corrected instrument View From Top (+Y)
source
sample
Silicon window
detector
SampleToSi (swd)
sample_detector_distance (sdd)
(0, 0, 0)
+ Z
+ X
Center of silicon window is at nominal position/origin (0, 0, 0)
Sample is at (0, 0, -swd)
Detector is at (0, 0, sdd - swd)

## Slide 8: Overwriting sample to Silicon window distance
source
sample
Silicon window
detector
Overwrite SampleToSi (swd’)
sample_detector_distance (sdd + (swd’ – swd))
(0, 0, 0)
+ Z
+ X
User overwrites sample to silicon window distance to swd’
Sample is moved to (0, 0, -swd’)
Detector is NOT moved
Sample to detector distance is changed to (sdd + swd’ – swd) with an offset (swd’ – swd), i.e., the offset of the overwritten value to EPICS value
(0, 0, -swd’)

## Slide 9: Overwriting sample to detector distance
source
sample
Silicon window
detector
Overwrite SampleToSi (swd’)
sample_detector_distance (sdd’)
(0, 0, 0)
+ Z
+ X
User overwrites sample to detector distance to sdd’
Sample is NOT moved
Detector is shifted by offset (sdd’ – sdd), such it is moved to (sdd ‘– swd)
Sample to detector distance is changed to the overwritten value sdd’
(0, 0, -swd)
(0, 0, sdd’ - swd)

## Slide 10: Summary
The sample detector distance specified by instrument scientist/user in EPICS can be changed by meta data
  sample to Silicon window position or/and
  sample to detector distance
If both sample to Silicon window and sample to detector distance is overwritten,
  sample is moved to new position by overwritten sample silicon window distance
  detector position is recalculated from new sample position and overwritten sample to detector distance

## Slide 11: Example
Data: /HFIR/CG2/IPTS-23801/nexus/CG2_7116.nxs.h5
Meta
  CG2:CS:SampleToSi = 88 mm
  sample_detector_distance = 6.88799993 m
Without any meta data overwriting
  Sample is located at (0, 0, -0.088)
  Detector is located at (0, 0, 6.800)
  Sample detector distance is 6.888 m
Overwrite sample to silicon window distance to 0 mm
  Sample is located at (0, 0, 0)
  Detector is located at (0, 0, 6.800)
  Sample detector distance is 6.800 m
Overwrite sample to detector is distance to 60.888 m
  Sample is located at (0, 0, -0.088)
  Detector is located at (0, 0, 60.800)
  Sample detector distance is 60.888 m

## Slide 12: Bio-SANS

## Slide 13: Instrument View From Top in the world of instrument scientists and EPICS
source
sample
Silicon window
main detector
SampleToSi
sample_detector_distance (sdd)
wing detector

## Slide 14: From the world of IS/EPICS to the world of drt-sans
In drt-sans, data from a NeXus file is loaded into a Mantid workspace
In a Mantid workspace, an instrument (object) is created for
  Neutron source position
  Sample position
  Detector pixels position on main detector
  Detector pixels positions on wing detector
Instrument geometry is set to a Cartesian coordinate system in drt-sans (Mantid)
  A fixed origin/nominal position set to the intersection of neutron beam and wing detector’s arm
  Sample can be moved
  Detector can be moved
  Source can be moved
  Silicon window CANNOT be moved and fixed at 71mm from nominal position (original) along neutron beam’s direction. Thus its position is (0, 0, -0.071) meter.

## Slide 15: From the world of IS/EPICS to the world of drt-sans (cont)
Meta data in the 3 worlds
Instrument scientists | EPICS | drt-sans
Sample to silicon window distance | SampleToSi | swd
Sample to main detector distance | sample_detector_distance | sdd

## Slide 16: drt-sans: Expected Instrument View From Top (+Y)
source
sample
Silicon window
main detector
SampleToSi
sample_detector_distance (sdd)
wing detector
(0, 0, 0)
71 mm
(0, 0, -(swd – 0.071 ) unit = meter
(0, 0, sdd - swd + 0.071 ) unit = meter
Nominal position(original) is the interaction of wing detector’s arm and neutron beam
Center of silicon window is at (0, 0, -0.071) (unit is meter), i.e, 71mm from origin
Sample is at (0, 0, -(swd – 0.071)) (unit is meter)
Detector is at (0, 0, sdd – swd + 0.071) (unit is meter)

## Slide 17: drt-sans: Instrument view from Top in Current Implementation (wrong)
source
sample
Silicon window
main detector
sample_detector_distance (sdd)
wing detector
(0, 0, 0)
(0, 0, sdd - swd + 0.071 ) unit = meter
Issue: sample always is positioned at (0, 0, 0). Sample to silicon distance is never considered
Is it critical?
Yes! Any offset to sample or main detector position can result in wrong detector pixels’ positions on wing detector
Cause: Mantid instrument definition file (IDF) does not consider meta data SampleToSi
Solution:
Move the sample toward source by swd – 71 mm to (0, 0, – swd + 0.071 meter)
Move the detector toward source/sample by swd - 71 mm to (0, 0, sdd – swd + 0.071 meter)
Such that the sample-detector distance is not changed from the value recorded in EPICS
Then the instrument will be corrected to what is shown in previous slide.

## Slide 18: drt-sans: Corrected instrument View From Top (+Y)
source
sample
Silicon window
main detector
SampleToSi
sample_detector_distance (sdd)
wing detector
(0, 0, 0)
71 mm
(0, 0, -(swd – 0.071 ) unit = meter
(0, 0, sdd - swd + 0.071 ) unit = meter
Nominal position(original) is the interaction of wing detector’s arm and neutron beam
Center of silicon window is at (0, 0, -0.071) (unit is meter), i.e, 71mm from origin
Sample is at (0, 0, -(swd – 0.071)) (unit is meter)
Detector is at (0, 0, sdd – swd + 0.071) (unit is meter)

## Slide 19: Overwriting sample to Silicon window distance
source
sample
Silicon window
main detector
SampleToSi = swd’
sample_detector_distance (sdd)
wing detector
(0, 0, 0)
71 mm
(0, 0, -(swd’ – 0.071 ) unit = meter
(0, 0, sdd - swd + 0.071 ) unit = meter
User overwrites sample to silicon window distance to swd’
Sample is moved to (0, 0, -swd’ + 0.071) unit =meter
Main detector is NOT moved
Sample to detector distance is changed to (sdd + swd’ - swd) with an offset (swd’ – swd), i.e., the offset of the overwritten value to EPICS value
Wing detector is NOT moved

## Slide 20: Overwriting sample to detector distance
source
sample
Silicon window
main detector
SampleToSi
sample_detector_distance (sdd)
wing detector
(0, 0, 0)
71 mm
(0, 0, -(swd – 0.071 ) unit = meter
(0, 0, sdd’ - swd + 0.071 ) unit = meter
User overwrites sample to detector distance to sdd’
Sample is NOT moved but still at (0, 0, – swd + 0.071), unit = meter
Detector is shifted by offset (sdd’ – sdd), such it is moved to (sdd’ – swd + 0.071)
Sample to detector distance is changed to the overwritten value sdd’

## Slide 21: Summary
The sample detector distance specified by instrument scientist/user in EPICS can be changed by meta data
  sample to Silicon window position or/and
  sample to detector distance
If both sample to Silicon window and sample to detector distance is overwritten,
  sample is moved to new position by overwritten sample silicon window distance
  detector position is recalculated from new sample position and overwritten sample to detector distance

## Slide 22: Example
Data: /HFIR/CG3/IPTS-23782/nexus/CG3_4829.nxs.h5
Meta
  CG3:CS:SampleToSi = 71 mm
  sample_detector_distance = 7.00 m
Without any meta data overwriting
  Sample is located at (0, 0, 0)
  Detector is located at (0, 0, 7.00)
  Sample detector distance is 7.00 m
Overwrite sample to silicon window distance to 82.3 mm, i.e., shift 12.3 mm from position recorded by EPICS
  Sample is located at (0, 0, 0.0123)
  Detector is located at (0, 0, 7.00)
  Sample detector distance is 7.0123 m
Overwrite sample to detector distance to 7.1234 m
  Sample is located at (0, 0, 0)
  Detector is located at (0, 0, 7.1234)
  Sample detector distance is 7.1234 m